## Slide 1: Low Impedance Measurement Using Low Cost SDR VNA
Shrinithi Venkatesan
Adinath Phene
February 13, 2024

## Slide 2
Need for Low Impedance Measurement
Two Port Measurement Method
Lowest Impedance measured by the SDR VNA
RLC Characteristics Trends of Circuit Elements
Takeaways
Agenda

## Slide 3
PDN components have target impedances in the mOhm range.
Not practical to measure Z < 0.1 Ohms using return loss 1-port VNA.
Real world limitations of Signal to Noise ratio & fixturing reproducibility.
2- Port method is a low-cost method to measure impedances even < 0.1 ohms.
Need for Low Impedance Measurement
Low Z
Measurement
Vias
Solder
Blobs
MLCC
SMT
Resistors
PDN
VRM

## Slide 4
4 – Point Kelvin Method for Low Z Measurement
Measurement Set up for Calculating the Resistance.
Rated Resistors (in Ohm)​ | Measured Resistance (Uncertainty ± 0.002) (in Ohm)
10 | 10.006
1 | 1.003
0.50 | 0.505
Measured resistors are within 1% tolerance of the rated values.
Next, we will correlate the measured & simulated impedance.

## Slide 5
Low Impedance Measurement – 1 & 2 Port Method
2- Port Method (Z < 0.1 ohms)
1- Port Method (Z > 0.1 ohms)

## Slide 6
Open, short & 50-ohm load plotted on a Smith Chart.​
SDR VNA is calibrated.
SDR VNA Calibration – Open, Short, Load & Through
Thru - Insertion Loss is 0dB (100% Transmission)
Noise Floor is the lowest impedance measured by the SDR VNA.
Open – S21 (Noise Floor)
Through – S21 (Insertion Loss)
Calibration
Open
Short
50 Ohm Load

## Slide 7
Correlation in Measured & Simulated Resistor Impedance
1 Ohm Resistor
Breakout Board
Plotting Measured Impedance in ADS

## Slide 8
Measured & Simulated 1 Ohm Resistor Impedance
This simple RL model is an excellent model for the impedance behavior of this real resistor from low to high frequency.
Measured & Simulated Impedance correlate accurately.
Measured Impedance - 1 Ohm Resistor
Simulated Impedance - 1 Ohm Resistor
R
ESL
Rated Resistor (in Ohm)​ | Modeled R (in Ohm) | Modeled ESL (in nH)
1.00 | 1.00 | 1.47

## Slide 9
R & L Characteristics of SMT Resistors
Rated Resistors (in Ohm)​ | ModeledR (in Ohm) | ModeledESL (in nH)
1​0 | 10.00 | 1.47
1 | 1.00 | 1.47
0.5 | 0.50 | 1.47
Zero ​Ohm Jumper | 0.011 | 1.45
10 Ohm
1 Ohm
0.5 Ohm
Zero Ohm
Similar ESL values due to identical footprint.
Zero Ohm Jumper has the lowest resistance.
Measured Impedance All resistors
Simulated Impedance &
10 Ohm
1 Ohm
0.5 Ohm
Zero Ohm Jumper

## Slide 10
Lowest Impedance Measured by SDR VNA
Resistor Structures | Modeled R (in mOhm) | Modeled ESL (in nH)
Zero Ohm Jumper | 11 | 1.45
Solder Blob | 0.55 | 0.45
Noise Floor | 0.15 | 0.00025
Noise Floor Resistance: 0.15 mohm Noise Floor Inductance: 0.25 pH
Solder Blob
(Made wide & thick to have
Low R & L)
Lowest Impedance measured by the SDR VNA is the Noise Floor.
Zero Ohm Jumper
Noise Floor
Measured Impedance
&
Simulated Impedance

## Slide 11
The 2-port method can eliminate the artifacts associated with contact impedance of the probes or fixturing to the DUT.
2- Port Method is a low-cost and reliable method for Low Z Measurement.
First order RL models can be used to hack the impedances measured with the 2-port method to correlate measured & simulated data.
SDR VNA can measure resistances & inductances in the mOhm and pH range respectively.
ESL of components depends on their footprint.
Takeaways

## Slide 12
Questions?

## Slide 13
RLC Equivalent Z for an SMT MLCC
15 uF SMT MLCC
Breakout Board
Plotting Measured Impedance in Qucs

## Slide 14
Reduced Inductance of MLCC – X2Y
Board Structure | ESR (in mOhm) | ESL (in nH) | C (in uF)
Multiple Vias | 16.3 | 1.225 | 0.20
X2Y | 9.63 | 0.675 | 0.368
Simulated Impedance
Measured Impedance
Multiple Vias
X2Y Capacitor
X2Y Capacitor
Multiple Vias
X2Y capacitors have a much lower inductance than Multiple Vias structure.
There is a reduction of a factor of 2 in ESL.
This translates into a reduction in the impedance at high frequency by a factor of 2 and could enable a parts count reduction by a factor of 2 as well.

## Slide 15
RLC Equivalent Z for Electrolytic & Tantalum Capacitors
Rated Capacitance (in uF)​ | ESR (in Ohm) | ESL (in nH) | C (in uF)
15 (SMT MLCC) | 0.00366 | 0.75 | 13.2
0.47 (Electrolytic) | 4 | 8.4 | 0.39
2.2 (Tantalum) | 0.57 | 7.70 | 2.28
This is the best match we can get with a first order simulated RLC model.
ESR for the SMT MLCC is much lower than the Electrolytic & Tantalum Capacitors.
Electrolytic & Tantalum Capacitors have a much higher ESL due to the leads.
Measured Impedance
Simulated Impedance
Electrolytic Cap
SMT MLCC
Tantalum Cap
Tantalum Capacitor
Electrolytic Capacitor

## Slide 16
RLC Equivalent Z for an SMT MLCC
This simple RLC model is good enough to match the impedance behavior of the real SMT MLCC component.
Rated Capacitance (in uF) | ESR (in mOhm) | ESL (in nH) | C (in uF)
1​5 | 3.6 | 0.75 | 13.2
ESL for SMT MLCC is extremely low & this is the best match we can get using a first order model.
Simulated Impedance – SMT MLCC
Measured Impedance – SMT MLCC
ESL
ESR
C

## Slide 17
ESL of 0.22 uF MLCC with different Via positions
Via Position | ESR (in mOhm) | ESL (in nH) | C (in uF)
20 mils from Pad | 38.8 | 4.05 | 0.20
5 mils from Pad | 22.3 | 2.15 | 0.20
In the Pad | 17 | 1.7 | 0.20
Multiple Vias | 16.3 | 1.225 | 0.20
Vias
20 mils from Pad
Vias
5 mils from Pad
Multiple
Vias
Vias
in the Pad
Simulated Impedance
Vias 20 mils from Pad
Vias 5 mils from Pad
Vias in the Pad
Multiple Vias
ESL is not an intrinsic measure of the capacitor itself as it is strongly dependent on how the capacitor is mounted to the planes.
Measured Impedance
All Structures